## Slide 1: Minimally Invasive Annuloplasty Technology for Tricuspid and Mitral Valve Repair
Mathew Williams, MD
NYU Langone Medical Center
Thursday, May 23, 2019

## Slide 2: Potential conflicts of interest (1)
Speaker's name : Mathew Williams
☑ I have the following potential conflicts of interest to declare:
Receipt of grants / research supports: Abbott, Edwards Lifesciences, HLT, Livanova, Medtronic
Stock shareholder: Micro Interventional Devices, Thubrikar Valve

## Slide 3
  Compliant anchors “conform” to beating heart
  Retracting barbs facilitate tissue penetration and anchor placement
  Expanding barbs resist pull out up to a 4N load
  Biocompatibile
  Polypropylene used in millions of implants
Ultra low mass implant
  Complies with natural cardiac function
Clinically Validated
  Over 45 patients
  No Clinical Adverse Events
Anchor platform engineered for catheter-based percutaneous fixation
Investigational Device ONLY.
PolyCorTM - Anchor Engineered for Cardiac Fixation

## Slide 4
MIA Implant Pre/Clinical Safety and Performance Data
High Speed Delivery Systems:
Implant Deployed in 0.002 seconds
Investigational Device ONLY.
High Speed, Delivery Systems

## Slide 5
14.74 mm
Current Anchor
Enhanced Anchor
10.24 mm
PolyCor - Anchor Engineered for Cardiac Fixation

## Slide 6
MIA - Simple, Safe and Secure Technology
Applying minimally invasive surgical and interventional knowledge to achieve surgical efficacy with percutaneous safety.
Investigational Device ONLY.

## Slide 7
MIA Tricuspid System, Implant and Stabilizing Unit
Representative 3D Echo of implant locations between AP and PS commissures (area of annular dilatation)
Implant
Delivery Catheter
Steerable Guiding Sheath
Stabilization Unit
Confidential – For Internal Use Only
Investigational Device ONLY.

## Slide 8
Technology Replicates Surgical Procedures
Standard Surgical Bicuspidization
MIA Replicates Clinical Surgical Technique
Percutaneous Tricuspid MIA Procedure
Photos courtesy of Professor Kestutis Rucinskas, MD, Chief of Cardiac Surgery
Vilnius University Hospital Santariskiu Clinics
Confidential – For Internal Use Only
Investigational Device ONLY.

## Slide 9
45% Reduction in annular area
STTAR Feasibility Arm - Surgical Results
Results at 12 Month Follow Up
Investigational Device ONLY.

## Slide 10
60 Patient; 6 Center CE Mark Study
  Vilnius University Hospital, Santaros Klinikos, Vilnius, Lithuania
  Kaunas University Hospital, Kaunas, Lithuania
  Semmelweis University, Budapest, Hungary
  Cardinal Wyszynski National Institute of Cardiology, Warsaw, Poland
  Pauls Stradins Clinical University Hospital, Riga, Latvia
  Pending - Aarhus University Hospital Denmark, Aarhus, Denmark
Image- Guided Percutaneous Arm
  Pre procedure CT and Echo
  Fluoroscopy, 2D and 3D Echocardiography
End Points:
  Safety:
    Major Adverse Events
  Performance:
    Technical Success Rate
    Reduction in Tricuspid Regurgitation/EROA
    Six Minute Walk Test
    QOL
Investigational Device ONLY.
FIH Vilnius, Lithuania (October 2018)
STTAR: Study of Transcatheter Tricsupid Annular Repair

## Slide 11
Inclusion Criteria:
  Age ≥18
  New York Heart Association (NYHA) Class II-IV
  Left ventricular ejection fraction (LVEF) ≥30%
  Tricuspid valve annular diameter ≥ 40 mm (or 21 mm/m2) 3-4 + FTR
Exclusion Criteria:
  Previous tricuspid valve repair or replacement
  Life expectancy of less than 12-months
  Pulmonary hypertension with PA mean 2/3 rd MAP
  Severe right heart dysfunction
Investigational Device ONLY
Study Design: Prospective non-randomized safety and performance study
Variables | Mean (Range)
Age, years | 74(61-85)
Euroscore II | 4.75(1-14.2)
Gender,female | 7(58%)
NYHA II | 2(20%)
NYHA III | 7(70%)
NYHA IV | 3(30%)
Atrial fibrillation | 6(50%)
Pulmonary hypertension | 6(50%)
Pacemaker | 3(30%)
Patient Criteria n = 10
Investigational Device ONLY.
STTAR: Study of Transcatheter Tricsupid Annular Repair

## Slide 12
FIH Transcatheter Procedure
Video courtesy of Professor Kestutis Rucinskas, MD, Chief of Cardiac Surgery; Vilnius University Hospital Santariskiu Clinics
Fluoroscopy Pre-Plication – Annular Area 14.2cm2
Fluoroscopy Post-Plication – Annular Area 7.6cm2
STTAR: Study of Transcatheter Tricsupid Annular Repair
Notes: We used trans-jugular approach to reach right atrium with steerable guiding sheeth. Further delivery catheter insertion and navigation with TEE. After identification of antero-posterior commissure we stabilize delivery system by deploying stabilization pins which also help us navigate close to right coronary artery and ensure we do not deploy into unnecesary anatomical structures. Anchors deployed into the posterior portion of the annulus. Then we plicate the tissue by approximating the anchors and securing the bicuspidization with a suture lock.

## Slide 13: STTAR: Percutaneous Results
*Mann-Whitney U test
*
Pre Plication
Post Plication
Investigational Device ONLY
33%
Acute results n = 10 (p = 0.0001)
26.4%
3 Month Follow Up n = 4 (p = 0.007)
26.4%
Investigational Device ONLY.
STTAR: Study of Transcatheter Tricsupid Annular Repair

## Slide 14
STTAR Clinical Case – Budapest, Hungary
Pre Plication on Fluoroscopy – 14.5 cm2
Post Plication on Fluoroscopy – 6.94 cm2
52% Reduction in Annular Area
Radiopaque anchors on echo
Investigational Device ONLY.
  Photos courtesy of the Semmelweis University, Budapest, Hungary
STTAR: Study of Transcatheter Tricsupid Annular Repair

## Slide 15
Severe TR to Trace TR
Color Doppler Prior to MIA Plication
Color Doppler Post MIA Plication
STTAR Clinical Case – Budapest, Hungary
Investigational Device ONLY.
  Images courtesy of the Semmelweis University, Budapest, Hungary
STTAR: Study of Transcatheter Tricsupid Annular Repair

## Slide 16
3D Echo Prior to MIA Plication
3D Echo Post MIA Plication
Investigational Device ONLY.
  Images courtesy of the Semmelweis University, Budapest, Hungary
STTAR Clinical Case – Budapest, Hungary
STTAR: Study of Transcatheter Tricsupid Annular Repair

## Slide 17
Mitral Valve before Implantation
Mitral Valve after second plication
Array 1 Locked
Arrays 1 & 2 Locked
Annulus Pre-procedure
Preclinical results of a Kay Procedure Annuloplasty
Preclinical results of a Continuous Ring Annuloplasty
MIA –Mitral Repair Program Working Towards FIH
Investigational Device ONLY.

## Slide 18
  A 60 patient STTAR CE Mark study is ongoing in Europe
  Safety profile is high – no serious device-related adverse events peri-procedurally or during follow-up
  Significant and durable reductions in tricuspid regurgitation and annular dimensions have been achieved
  Technology treats a wide range of functional TR patients with minimal exclusion criteria
  Procedure is simple and the learning curve appears to be short, with procedure times as low as 70 minutes in early transcatheter experience
A simple, safe and secure transcatheter repair technology
Replicates surgical procedures on the end of a catheter
Investigational Device ONLY.